## Sheet1
CHANGE LOG
v7.2 | 1. Add an event ofEVENT_FENOM_SCORE to notify POC APP | 2. replace battery with the serial number(4 bytes) in Device info
v7.1 | Add an event of EVENT_T2_REAL_TIME to notify POC APP
v7 | Add Breath Gauge Pressure, NO Counts, Sample Mass Flow to Breath Maneuver characteristic.
v6.1 | Add Descriptor attribute to characteristic
v6 | Delete Get Log characteristic
v5 | Define Status Code for Breath Maneuver char. Define 0ppb test number in Breath Test characteristic. Define Training Mode characteristic. Define Get Log characteristic
v4 | Update Breath Maneuver char. Add examples
v3 | Define new characteristics: Environmental Info Breath Test Breath Maneuver
v2 | Define Data Command Service and Device Info characteristic
Data Command Service
| Device Info characteristic | Byte position | Definition | Range | Unit | Example
| Property: read | 0 | Device status | 1: ready 0:not ready | N/A | 0x01 | 1
| Descriptor: "Device Info" | 1 | FW major version | [0x00; 0xFF] | N/A | 0x00 | 0
|  | 2 | FW minor version | [0x00; 0xFF] | N/A | 0x00 | 0
|  | 3 | FW build number | [0x00; 0xFF] | N/A | 0x01 | 1
|  | 4 | Sensor Exp Date (Month) | [0x01; 0x0C] | N/A | 0x0C | 12
|  | 5 | Sensor Exp Date (Day) | [0x01; 0x1F] | N/A | 0x1F | 31
|  | 6 | Sensor Exp Date (Year) | 0x02 | N/A | 0x08 | 2099
|  | 7 | Sensor Exp Date (Year) | 0x00 | N/A | 0x33
|  | 8 | serial number of F150 | [0x00; 0xffff(65535)] | % | 0x64 | 100
| Environmental Info characteristic | Byte position | Definition | Range | Unit | Example
| Property: read | 0 | Temperature | [0x00; 0x58] | Celsius | 0x19 | 25
| Descriptor: "Environmental Info" | 1 | Humidity | [0x14; 0x5A] | % | 0x32 | 50
|  | 2 | Pressure | [0x78; 0x32] | hPa | 0x65 | 101*10=1010
| Breath Test characteristic | Byte position | Definition | Range | Unit
| Property: write without response Descriptor: "Breath Test" | 0 | Test number | 0: stop test 1: start 6-sec test 2: start 10-sec test 3: start 0ppb test | N/A | Note: for 0ppb test result, only NO score is valid when reading Breath Maneuver characteristic.
Note: READ Breath Maneuver characteristic for test results
Status Code
| Breath Maneuver characteristic | Byte position | Definition | Range | Unit | Example |  | 0x00 | Test performed OK
| Property: read | 0 | Test Number MSB | [0x0000; 0xFFFF] | N/A | 0x1B | 7020 | 0x01 | Error: low battery
| Descriptor: "Breath Maneuver" | 1 | Test Number LSB |  |  | 0x6C |  | 0x02 | Error: breath flow rate > 4.4 LPM anytime during T2 | see FSRS-5.9.6
|  | 2 | Time Remaining | [0x00; 0xFF] | second | 0x03 | 3 | 0x03 | Error: breath flow rate < 1 LPM anytime during T2 | see FSRS-5.9.6
|  | 3 | Temperature | [0x00; 0x58] | Celsius | 0x19 | 25 | 0x04 | Error: breath flow rate outside [2.5; 3.5] LPM for > 3 seconds during washout | see FSRS-3.2.2
|  | 4 | Pressure | [0x4C; 0x6E] | x10 hPa | 0x65 | 101*10=1010 | 0x05 | Error: breath flow rate outside [2.7; 3.3] LPM for > 3 seconds after washout | see FSRS-3.2.3
|  | 5 | Breath Flow | [0x1B; 0x21] | /10 LPM | 0x1D | 29/10=2.9 | 0x06 | Calculating test results | see FSRS-5.1.6
|  | 6 | NO Score MSB | [0x0005; 0x012C] | ppb | 0x01 | 299 | 0x07 | Error: temperature too high (ambient temperature > 40 C)
|  | 7 | NO Score LSB |  |  | 0x2B |  | 0x08 | Error: temperature too low (ambient temperature < 15 C)
|  | 8 | Analysis Time Left | [0x00; 0xFF] | second | 0x03 | 3 | 0x09 | Error: too dry (humidity < 20% relative humidity)
|  | 9 | Status Code | see Status Code | N/A | 0x00 | 0 | 0x0A | Error: too humid (humidity > 90% relative humidity)
|  | 10 | Breath Gauge Pressure | [5; 68] or [0x05; 0x44] | /10 kPa | 0x09 | 9/10=0.9 | 0x0B | Error: pressure low (atmospheric pressure < 760 hPa)
|  | 11 | NO Counts MSB | [0; 32768] or [0x0000; 0x8000] | N/A | 0x20 | 8226 | 0x0C | Error: pressure high (atmospheric pressure > 1100 hPa)
|  | 12 | NO Counts LSB |  | N/A | 0x22
|  | 13 | Sample Mass Flow | [100; 250] or [0x64; 0xFA] | sccm | 0x64 | 100
| Training Mode characteristic | Byte position | Definition | Range | Unit
| Properties: read, write without response Descriptor: "Training Mode" | 0 | Mode # for target flow rate on WRITE | 0: stop training mode. 1: pratice breathing about 2.5 LPM into the device 2: pratice breathing about 3.0 LPM into the device. 3: pratice breathing about 3.5 LPM into the device. | LPM on READ | Note: when reading, the actual flow rate is returned during training.
| Event name | Initiator task | receiving task | transmitting data
1 | EVENT_T2_REAL_TIME(1<<13) | Central processing task | RF_Thread_Process | breath flow(0~500) | 500 means 5LPM
2 | EVENT_FENOM_SCORE(1 << 29) | Central processing task | RF_Thread_Process | fenom Score(0~500) | unit:ppb